## Slide 1
MODULE 3: PROGRAMMING FOR BIOLOGICAL DATA
Hypothesis Testing in the Clinical Laboratory
Session 5: T-tests, Normality, and P-values
Instructor: Dr. Haogao Gu | Date: January 19, 2026

## Slide 2
Distribution
Understand the Normal Distribution and its critical role in reference intervals.
Normality
Test data for normality using shapiro.test() and Q-Q plots.
T-Tests
Perform one-sample, paired, and two-sample t-tests in R.
Interpretation
Interpret p-values correctly in clinical decision-making contexts.
Learning Objectives

## Slide 3
The Questions We Ask
The Objective Answer
Hypothesis testing provides a rigorous statistical framework to answer these questions objectively, moving beyond "it looks different" to "it is statistically different."
Is this new reagent lot equivalent to the old one?
Is our analyzer calibrated correctly?
Do viral loads differ between patient groups?
Why Hypothesis Testing?

## Slide 4
Part 1: The Normal Distribution
Foundations of Clinical Statistics
Duration: 40 minutes

## Slide 5
The Normal (Gaussian) Distribution is the foundation of many laboratory practices.
Application | How Normal Distribution is Used
Reference Intervals | The standard "Mean ± 2SD" calculation covers 95% of a healthy population only if the data is normally distributed.
Quality Control | Levey-Jennings charts rely on the assumption that control values follow a normal distribution.
Statistical Tests | Most parametric tests (like t-tests and ANOVA) assume normality to be valid.
Why "Normal" Matters in the Lab

## Slide 6
Visualizing Hemoglobin
Imagine measuring hemoglobin in 1000 healthy adults:
•
Most values cluster around the center (the mean).
•
Fewer values exist at the extremes (very low or very high).
•
The shape is perfectly symmetric.
The Bell Curve

## Slide 7
Standard Deviation Rules
Clinical Application
This captures the central 95% of the healthy population.
•
Mean ± 1SD: Covers ~68% of values
•
Mean ± 2SD: Covers ~95% of values
•
Mean ± 3SD: Covers ~99.7% of values
Key Properties

## Slide 8
Visualizing the Normal Distribution in R

## Slide 9
R Code
Expected Output
Based on our simulated data (Mean=14, SD=1.5):
This interval would theoretically contain 95% of healthy individuals.
Reference Interval: 11.0 - 17.0 g/dL
Calculating Reference Intervals

## Slide 10
Before using parametric tests (like t-tests), you must verify your data is approximately normal.
Method | Type | Tool in R | What to look for
Q-Q Plot | Visual | qqnorm(), qqline() | Points falling along the diagonal line.
Shapiro-Wilk | Statistical | shapiro.test() | P-value > 0.05 (Fail to reject null).
Testing for Normality

## Slide 11
How to Read It
A Q-Q (Quantile-Quantile) plot compares your data against a theoretical normal distribution.
Normal: Points follow the straight diagonal line closely.
Skewed/Non-normal: Points curve away from the line (S-shape or banana shape).
Q-Q Plot: Visual Normality Check

## Slide 12
Code
Visual Inspection
The code on the left generates two plots side-by-side.
Notice how the Normal Data hugs the red line, while the Skewed Data (exponential distribution) deviates significantly at the ends.
Live Demo: Q-Q Plots in R

## Slide 13
The Hypothesis
Interpretation
•
H₀ (Null): Data is normally distributed.
•
H₁ (Alternative): Data is NOT normally distributed.
p ≥ 0.05: Fail to reject H₀ → Data may be normal.
p < 0.05: Reject H₀ → Data is NOT normal.
Shapiro-Wilk Test

## Slide 14
1
Test First
Always test reference population data for normality before calculating intervals.
2
Non-Parametric
If data is not normal (and cannot be transformed), use non-parametric methods (percentiles) instead of Mean ± 2SD.
3
Transform
Consider data transformation (e.g., Log10) to normalize skewed distributions like viral loads.
Clinical Relevance: CLSI C28-A3

## Slide 15
Viral loads are typically right-skewed. Log-transformation often normalizes them, allowing the use of parametric tests.
Log Transformation for Skewed Data

## Slide 16
Part 2: T-Tests — Comparing Means
Statistical Comparisons in the Lab
Duration: 40 minutes

## Slide 17
Courtroom Concept | Hypothesis Testing Equivalent
Innocent until proven guilty | H₀: No difference (Null Hypothesis is assumed true initially)
Evidence must be convincing | P-value measures the strength of evidence against H₀
"Beyond reasonable doubt" | p < 0.05 → Reject H₀ (Statistically Significant)
"Not enough evidence" ≠ Innocent | p ≥ 0.05 → Fail to reject H₀ (Not proven equal, just no difference found)
The Logic: A Courtroom Analogy

## Slide 18
"The probability of obtaining results at least as extreme as the observed results, assuming that the null hypothesis is correct."
p < 0.05
Statistically Significant
Strong evidence against H₀.
p ≥ 0.05
Not Significant
Insufficient evidence to reject H₀.
Understanding the P-value

## Slide 19
Scenario | Test Type | R Function
Compare a sample mean to a known standard/target. | One-sample t-test | t.test(x, mu = value)
Compare same subjects under two conditions (e.g., Method A vs B). | Paired t-test | t.test(x, y, paired = TRUE)
Compare two independent groups (e.g., Delta vs Omicron patients). | Two-sample t-test | t.test(x, y)
Which T-Test to Use?

## Slide 20
Clinical Scenario
You receive a new lot of QC material.
The Question
Does your analyzer measure at the target?
H₀: Mean = 100 mg/dL
H₁: Mean ≠ 100 mg/dL
•
Manufacturer states target glucose = 100 mg/dL.
•
You run 20 replicates on your analyzer.
Scenario 1: One-Sample T-Test

## Slide 21
One-Sample T-Test in R
qc_glucose <- c(101, 103, 99, 102, 100, 104, 98, 103, 101, 102, 100, 105, 99, 101, 103, 102, 100, 104, 101, 103) target <- 100 # Perform t-test t.test(qc_glucose, mu = target) # Output: # t = 3.25, df = 19, p-value = 0.004 # mean of x: 101.55 # 95% CI: [100.5, 102.3]

## Slide 22
P-value
0.004
Since p < 0.05, we reject H₀. There is a statistically significant difference.
Mean Estimate
101.55
Our measured mean is ~1.55 mg/dL higher than the target of 100.
Decision
Bias Detected
The analyzer has a significant positive bias.
Interpreting the Output

## Slide 23
The Setup
Validating a new reagent (Method B) against the current one (Method A).
Why "Paired"?
We analyze the differences (A - B) for each patient.
If Mean Difference ≈ 0 → Methods are equivalent.
If Mean Difference ≠ 0 → Systematic bias exists.
•
Measure 20 patients with BOTH methods.
•
Samples are paired because measurements come from the same patient.
Scenario 2: Paired T-Test (Method Comparison)

## Slide 24
Result: A statistically significant bias of 1.5 mg/dL exists between the methods.
Paired T-Test in R

## Slide 25
The Bland-Altman (Difference) plot is the standard for method comparison.
•
X-axis: Mean of the two methods ((A+B)/2).
•
Y-axis: Difference between methods (A-B).
•
Horizontal Lines: Mean Bias (solid) and 95% Limits of Agreement (dashed).
Visualizing Difference: Bland-Altman Plot

## Slide 26
Statistical vs. Clinical Significance
Critical Concept!
Even if p < 0.05 (statistically significant), you must ask:
"Is the difference clinically meaningful?"
A 1.5 mg/dL bias in glucose might be statistically significant but is likely clinically irrelevant given biological variation.

## Slide 27
The Logic
Compare the observed bias against Total Allowable Error (TAE) specifications.
For Glucose, TAE is typically ±6 mg/dL or 10%.
R Implementation
Making a Clinical Decision

## Slide 28
Clinical Scenario
Do COVID-19 patients infected with Delta variant have different viral loads than Omicron?
•
Two independent groups of patients.
•
Different sample sizes possible.
Scenario 3: Two-Sample (Independent) T-Test

## Slide 29
Results
Interpretation
Recall that PCR Ct values are logarithmic.
A difference of ~3.3 Ct represents a 10-fold difference in viral copies.
Conclusion: Delta patients had significantly higher viral loads (~10x).
•
Delta Mean Ct: 18.6
•
Omicron Mean Ct: 21.8
•
Difference: 3.2 cycles
Interpreting Variant Comparison

## Slide 30
1
Normality
Use Shapiro-Wilk or Q-Q plots. If failed, use Wilcoxon (non-parametric).
2
Independence
Samples must be independent (or properly paired for paired tests).
3
Variance
For two-sample tests, check equal variance. If unequal, use Welch's t-test (default in R).
Before ANY T-Test: Check Assumptions

## Slide 31
Scenario | Test | Key Assumption
Sample vs. Known Value | One-sample | Sample is normal
Same subjects, 2 conditions | Paired | Differences are normal
Two independent groups | Two-sample | Both groups normal
Always: Check normality first!
Summary: Which T-Test?

## Slide 32
Tutorial 5
Method Validation Study: Practical Application
Open your R environment

## Slide 33
Role
You are a Medical Laboratory Technologist validating a new glucose reagent lot.
Task
•
Data: 20 patient samples measured by Method A (Current) and Method B (New).
•
Goal: Determine if there is systematic bias.
•
Question: "Is the new lot statistically and clinically equivalent?"
Tutorial Scenario

## Slide 34
Steps 1-2
Load Data & Calculate Differences
Step 3
Test Normality (Histogram/Shapiro)
Step 4
Perform Paired T-Test
Step 5-6
Clinical Significance & Bland-Altman Plot
What You'll Practice

## Slide 35
Load Data
Calculate Differences
Step 1 & 2: Load and Difference

## Slide 36
Before the t-test, verify the differences are normal.
Rule: If p ≥ 0.05, the data is normal, and we can proceed with the paired t-test.
Step 3: Test Normality

## Slide 37
Find these values in output:
•
t statistic: Magnitude of difference relative to variance.
•
p-value: Significance of the difference.
•
95% CI: The range of the true bias.
•
mean of differences: The observed bias.
Step 4: Perform Paired T-Test

## Slide 38
Check Clinical Limits
Bland-Altman Plot
Step 5 & 6: Significance & Plot

## Slide 39
Write a 3-5 sentence summary including:
1.
Number of samples compared (n=20).
2.
Mean bias and direction (Method A vs B).
3.
Statistical significance (p-value).
4.
Clinical acceptability (vs. ±6 mg/dL).
5.
Final Recommendation: Approve or Reject?
Challenge: Validation Report