## Slide 1
MODULE 3: PROGRAMMING FOR BIOLOGICAL DATA
Session 6: ANOVA and Linear Regression
Comparing Multiple Groups & Building Calibration Curves
Instructor: Dr. Haogao Gu | Date: January 21, 2026

## Slide 2
ANOVA
Compare means across multiple groups using Analysis of Variance.
Post-hoc
Identify exactly which groups differ using Tukey HSD.
Regression
Build linear regression models for standard calibration curves.
Validation
Assess model fit using R² and rigorous residual analysis.
Learning Objectives

## Slide 3
Method Applications
Method | Application
ANOVA | Compare 3+ analyzers, reagent lots, or operators.
Regression | Build calibration curves and verify assay linearity.
Today's Scenarios
•
Scenario 1: Are our three chemistry analyzers giving consistent results?
•
Scenario 2: Is our assay linear across the reportable range according to CLSI EP06?
Clinical Context

## Slide 4
Part 1: Analysis of Variance (ANOVA)
Comparing Means Across Multiple Groups

## Slide 5
The Right Test for the Job
Scenario | Statistical Test
Compare 2 groups | T-test
Compare 3+ groups | ANOVA
⚠️ Warning
Running multiple t-tests inflates the false positive rate!
•
Comparing 3 groups requires 3 separate t-tests.
•
This leads to a ~14% chance of finding a false difference.
•
ANOVA controls this error by testing all at once.
When to Use ANOVA

## Slide 6
The Setup
Your lab has three chemistry analyzers: A, B, and C.
You run the same glucose QC material 10 times on each instrument.
Question: Are all three analyzers calibrated correctly?
Preliminary Data
Analyzer | Mean Glucose | SD
A | 102.5 | 1.5
B | 98.2 | 1.4
C | 106.8 | 1.6
Scenario: Analyzer Comparison

## Slide 7
Visual inspection: The boxplot clearly suggests that the analyzers are producing different results, with C being notably higher.
Visualizing with Boxplots

## Slide 8
The F-Statistic
ANOVA compares between-group variance (differences between analyzers) to within-group variance (precision error).
Large F = Groups are different Small F = Groups are similar
How ANOVA Works

## Slide 9
Null Hypothesis (H₀)
All group means are equal.
Alternative Hypothesis (H₁)
At least one mean differs from the others.
Decision Rule: If p < 0.05, we reject H₀.
ANOVA Hypotheses

## Slide 10
R Code
Output
ANOVA in R

## Slide 11
Column | Meaning
Df | Degrees of freedom (Groups - 1)
Sum Sq | Sum of squares (total variance explained)
Mean Sq | Mean square = Sum Sq / Df
F value | Test statistic (Variance Ratio)
Pr(>F) | P-value (The probability of this result if H₀ is true)
p < 0.05 → At least one analyzer differs significantly.
Interpreting the ANOVA Table

## Slide 12
"ANOVA tells us that 'something' is different, but it doesn't tell us 'what' is different."
To find the specific culprit, we need a post-hoc test.
The Problem

## Slide 13
Purpose
Perform pairwise comparisons after a significant ANOVA to identify exactly which groups differ.
It controls for the increased error rate of multiple testing.
R Code
This compares:
•
A vs B
•
A vs C
•
B vs C
Tukey's Honest Significant Difference

## Slide 14
Column | Meaning
diff | Difference between group means.
lwr, upr | 95% Confidence Interval of the difference.
p adj | Adjusted p-value.
Reading Tukey Output

## Slide 15
Interpretation:
•
The lines represent Confidence Intervals.
•
If an interval crosses zero, the difference is NOT significant.
•
Here, all intervals are far from zero, indicating significant differences for all pairs.
Tukey Plot

## Slide 16
Based on the statistical findings, we take action:
Comparison | Difference | Significant? | Action
B vs A | -4.3 | Yes ⚠️ | Recalibrate B (reading low)
C vs A | +4.3 | Yes ⚠️ | Recalibrate C (reading high)
C vs B | +8.6 | Yes ⚠️ | Major discrepancy detected
Clinical Action Items

## Slide 17
1
Normality
Residuals should be normally distributed.
2
Homogeneity
Variances should be equal across groups.
3
Independence
Observations must be independent of each other.
Checking ANOVA Assumptions

## Slide 18
Part 2: Linear Regression & Calibration
Building the Standard Curve

## Slide 19
The Concept
Every quantitative assay uses a calibration curve to convert measured signals into concentrations.
We model this relationship using the linear equation.
Signal = Slope × Concentration + Intercept
The Foundation of Quantitative Testing

## Slide 20
Guideline Requirements
Standard for Lab Compliance
Test samples across the full reportable range.
Verify the linear relationship between signal and concentration.
Document recovery at each level.
Identify any saturation or non-linearity.
CLSI EP06: Linearity Verification

## Slide 21
Concentration (mg/dL) | Signal (OD) | Replicates
0 (blank) | 0.052 | 3
25 | 0.245 | 3
50 | 0.478 | 3
100 | 0.952 | 3
200 | 1.875 | 3
400 | 3.425 | 3
Linearity Verification Data

## Slide 22
Before fitting a model, always perform a visual check. Does the relationship look linear?
Visualizing Calibration Data

## Slide 23
Fitting the Model
We use the lm() function (linear model).
Viewing Results
The summary() function provides coefficients and fit statistics.
Linear Regression in R

## Slide 24
Term | Value | Meaning
Intercept | 0.052 | The background signal when concentration is 0.
Slope | 0.00872 | Signal increase per unit of concentration.
R² | 0.9987 | 99.87% of the variance is explained by the model.
Interpreting lm() Output

## Slide 25
R Code
Final Calibration Equation:
Getting the Equation

## Slide 26
This visualizes how well our calculated model fits the actual data points.
Adding the Regression Line

## Slide 27
Coefficient of Determination
Interpretation
R² Value | Meaning
0.99+ | Excellent fit (Standard for Clinical Labs)
0.95–0.99 | Good fit
< 0.90 | Poor fit — Investigate!
Understanding R²

## Slide 28
The Concept
Residual = Observed - Predicted
Residuals allow us to see patterns that R² might hide.
What to Look For
Random Scatter: Good model.
Curvature: Non-linearity (needs polynomial fit).
Funnel Shape: Heteroscedasticity (unequal variance).
Why Check Residuals?

## Slide 29
Creating a Residual Plot

## Slide 30
What R Generates
Plot | Checks For
Residuals vs Fitted | Linearity
Q-Q Plot | Normality of Residuals
Scale-Location | Homoscedasticity
Residuals vs Leverage | Influential Outliers
Built-in Diagnostic Plots

## Slide 31
R Code
Using the Calibration Curve

## Slide 32
Criterion | Typical Requirement
R² | ≥ 0.99
Residuals | Within ±10% of predicted value
Recovery | 95–105% at each concentration level
Note: Specific requirements vary by analyte and method SOP.
Linearity Acceptance Criteria

## Slide 33
Automated Verification Check
if (r_squared >= 0.99) { cat("✅ R² PASS") } else { cat("⚠️ R² FAIL") } if (max(abs(residuals)) <= 0.10 * predicted) { cat("✅ Residuals PASS") } else { cat("⚠️ Residuals FAIL") }

## Slide 34
Feature | ANOVA | Linear Regression
Purpose | Compare group means | Model relationships
Question | "Are groups different?" | "What is the equation?"
R Function | aov() | lm()
Key Statistics | F-value, p-value | R², slope, intercept
Lab Use | Analyzer/Lot Comparison | Calibration Curves
Summary: ANOVA vs. Regression

## Slide 35
ANOVA
Use for comparing 3+ groups. Always follow up with Tukey HSD.
Regression
Models y = mx + c. Use R² ≥ 0.99 as a quality standard.
Residuals
Essential for validating that your model assumptions are met.
Key Takeaways

## Slide 36
Tutorial 6
Calibration Verification in Practice
Open your R environment

## Slide 37
Tutorial 6: Calibration Verification
Building and Validating Standard Curves

## Slide 38
The Task
You are performing annual linearity verification following CLSI EP06 guidelines.
The Data
6 Concentration Levels
Each measured in triplicate.
•
Verify linear response across reportable range.
•
Document calibration equation.
•
Ensure R² ≥ 0.99.
Your Scenario

## Slide 39
Step | Skill | Time
1-2 | Load and explore data | 10 min
3 | Fit linear model | 15 min
4 | Create calibration plot | 15 min
5 | Analyze residuals | 15 min
6-7 | Calculate unknowns & Report | 15 min
What You'll Practice

## Slide 40
Load Data
Visualize
Steps 1-2: Load and Visualize

## Slide 41
The Magic Line
Extract Parameters
Step 3: Fit Linear Model

## Slide 42
Base R
ggplot2 (Publication Quality)
Step 4: Calibration Plot

## Slide 43
Why?
Residuals reveal problems R² misses:
Code
Goal: Random scatter around zero.
•
Curvature = Non-linearity
•
Funnel = Unequal variance
Step 5: Check Residuals

## Slide 44
Calculation
Step 6: Predict Unknowns

## Slide 45
Checklist
Example Conclusion
"Linearity verification PASSED. R² = 0.9987 exceeds acceptance criterion of 0.99. Assay is acceptable for clinical use."
•
Date and Analyst
•
Calibration Equation
•
R² Value (Pass/Fail)
•
Conclusion
Step 7: Verification Report

## Slide 46
Questions?
Session 6 Complete